## Version Control
Version | Date | Drafter/Reviewer | Remarks
1_0 | June 01, 2018 | WG3/PMTG

## Content description
This document contains checklists for SAE ARP6320 Procedure for the Continuous Sampling and Measurement of Non-Volatile Particulate Matter Emissions from Aircraft Turbine Engines Non-Volatile Particulate Matter Emissions from Aircraft Turbine Engines
| Sheet title | Description
Operational checklists used per engine test unless specified | Operational Procedure (OP) Calibration&Maintenance | Operational checklist for system calibration and maintenance items to be performed at different frequencies | Prior and Post engine series
Annually
Frequency as required
| OP System Installation | A list of consumables and items needed for installation including installation process
| OP Pre-engine start | Operation procedure and checks that need to be performed before starting the engine
| OP During engine running | A list of operational checks that need to be performed during the engine test; and nvPM sampling and measurement system procedure
| OP After engine shutdown | Operation procedure and checks that should be performed after engine shutdown
Parameter reporting checklist used per engine test report | Data System Parameters | List of parameters to be reported from engine test not included in nvPM transport loss calculation below
| REQ nvPM transport parameters | Required list of parameters to be reported for nvPM transport loss calculation
nvPM sampling and measurement system compliance performed routinely for a specific system | REQ Calibration&Maintenance | Compliance requirements for routine calibration, or change to instrument software/hardware
nvPM sampling and measurement system compliance performed once for a specific system | REQ General | System construction - list of requirements for compliance
| REQ System Section 1 | System construction - list of requirements for compliance
| REQ System Section 2 | System construction - list of requirements for compliance
| REQ System Section 3 | System construction - list of requirements for compliance
| REQ System Section 4 | System construction - list of requirements for compliance
| REQ System Section 5 | System construction - list of requirements for compliance
| REQ nvPMmi | List of requirements for instrument compliance
| REQ VPR | List of requirements for instrument compliance
| REQ nvPMni | List of requirements for instrument compliance
Cell color coding for Operational checklists and Data System Reporting
No data entry | Requirement | Recommendation or best practice

## OP Calibration&Maintenance
Test/Maintenance |  | Criteria | Result | Pass/Done |  | ARP 6320 requirement reference
Prior and Post engine test series:
| DF2 verification (at all settings) | Within ± 10% of annual calibration |  |  |  | 9.2.1 / A4.6
Annually (minimum):
| Empty and clean cyclone collection pot |  |  |  |  | 9.1.1
| Mass analyzer inlet flow rate calibration |  |  |  |  | 9.1.1
| VPR analyzer inlet flow rate calibration |  |  |  |  | 9.1.1
| Number analyzer inlet flow rate calibration |  |  |  |  | 9.1.1
| Make-up flow rate calibration |  |  |  |  | 9.1.1
CO2 analyzer flow rate calibration
| Optional: Validate system flow control by performing Section 3 total flow measurement | 25 ± 2 slpm
| Pressure transducers calibration | within 2% of nominal value |  |  |  | 9.1.1
| nvPMmi calibration (and maintenance if required) | Slope of the linear regression between mass instrument and EC determined by TOT |  |  | details in sheet "REQ Calibration&Maintenance" | 9.1.2 / 6.2
| nvPMni calibration (and maintenance if required) | Counting efficiency ≥50% at 10 nm |  |  |  | 9.1.4 / 7.2
|  | Counting efficiency ≥90% at 15 nm |  |  |  | 9.1.4 / 7.2
|  | Within ± 10% of Electrometer across measurement range |  |  |  | 9.1.4 / 7.2
|  | Pressure transducer within 2% of nominal value |  |  |  | 9.1.4 / 7.2
|  | Wick replacement |  |  |  | 9.1.4 / 7.2
| VPR calibration (and maintenance if required) | Minimum penetration efficiencies: 30% at 15nm 55% at 30nm 65% at 50nm 70% at 100nm |  |  |  | 9.1.3 / 7.1
|  | >99.5% removal of 30 nm Tetracontane aerosol |  |  |  | 9.1.3 / 7.1
|  | DF2 calibration (at all settings) |  |  |  | 9.1.3 / 7.1
| CO2, CO, HC & NOx analyzer calibration | ARP1256 audit requirements |  |  |  | 9.1.5 / ARP1256
Frequency as required:
| Diluter1 cleaning - needed when buildup of nvPM inside diluter sample orifice affects (increases) DF1 | Maintain DF1 within 8 to 14 |  |  |  | 9.1.1
| Optional: Aswell as monitoring increase of DF1, the procedure to determine Diluter1 operabilty with low inlet pressure at probe can be performed to assess whether Diluter1 cleaning is needed | Maintain DF1 within 8 to 14
| Replace Catalytic Stripper at manufacturer recommended interval |  |  |  |  | 9.1.3
| Replace Butanol in Number analyzer at manufacturer recommended interval |  |  |  |  | 9.1.4 / 7.2
| No data entry | Requirement | Recommendation or best practice

## OP System Installation
System Installation checklist |  | Yes (tick) | ARP 6320 requirement reference
Ensure provision of:
| DF1 diluent supply (Dry Air or Dry Nitrogen, <10ppm CO2, HEPA filtered) |  | 5.3.4
| DF2 diluent supply (Dry Air/N2, HEPA filtered, If VPR uses catalytic stripper diluent must contain minimum 10% Oxygen) |  | 7.1
Compressed air - dry and oil free (if required for valve operation and/or nvPMmi)
| High and low range CO2 span gas, specified by ARP1256 |  | 5.3.5 / 5.6.1
| CO, HC and NOx span gas (if needed), specified by ARP 1256 |  | 9.2.2
| Butanol |  | 7.2
For previously unused probe/rake for nvPM measurement - perform Diluter1 operability check at low sample pressure probe inlet conditions
Ensure Section 2 is within 8m of probe tip, connect to Collection Part outlet
Ensure Diluter1 vent line is short and with large diameter to minimise pressure drop. (This will help to prevent impact of DF1 operable range)
Ensure Section 4 inlet is within 25m of Section 2 outlet
Connect Section 3 to Section 2 and Section 4
Ensure Section 3 coiled bend radii is > 0.5 m |  |  | 5.4
Connect GL to Section 2 and Raw CO2 measurement system
Connect nvPMmi and VPR inlets to Section 4 outlets
Connect nvPMni inlet to VPR outlet
Install relevant system operation/control elements - including:
Power supply
Thermocouple cables
DAQ cables
Compressed gas tubing to - Diluter1, VPR, nvPMmi (if required) and CO2 analyzer
Ensure pump/diluter exhausts are safely vented to atmosphere
| No data entry | Requirement | Recommendation or best practice

## OP Pre-engine start
Test/Operation |  | Criteria | Result | Pass/Done | ARP6320 requirement reference
Sufficient compressed gas supply pressures to ensure completion of engine test:
DF1 Diluent
DF2 Diluent (VPR)
Compressed air (if required for nvPMmi)
CO2 span gases
Sufficient power available to all subsystems:
system heaters
pumps and flow controllers
Measurement analyzers (CO2, nvPMmi, VPR, nvPMni)
Data logging systems
System stabilized for a minimum of 1 hour |  |  |  |  | 9.2.2
System temperatures are within specified ranges.
Synchronization of data system logging clocks, if required
Data acquisition rate |  | ≥1 hz | _____ hz
Start system data logging
Spill and isolation valves are closed
Collection Part and GL leak check |  | ≤0.4 Litres per minute | ______ Litres in __ mins |  | 9.2.1 / A4.1
Collection Part and GL cleanliness (not required if performing simplified EI method) |  | HC < 1ppm or less than engine idle emission level | HC ______ ppm |  | 9.2.1 / A4.2
Gas analyzer operational checks (GL: CO2, CO, HC, NOx; Make-up flow CO2):
Sample, zero and span inlet flow rates
| Span and zero checks |  |  |  | 9.2.2
nvPM mass analyzer operational checks: |  |  |  |  | 9.2.2
Inlet flow rate
| Manufacturer recommended configuration is selected and operability checks (see relevant SOP) |  |  |  | 9.2.2
VPR analyzer operational checks: |  |  |  |  | 9.2.2
| VPR heated stage temperature | 623 +/-15 K
Inlet flow rate
| DF2 verification (all dilution settings that will be used during engine test) | ≤10% difference to calibration value | DF2 setting ____ _____ % different |  | 9.2.1 / A4.6
|  | ≤10% difference to calibration value | DF2 setting ____ _____ % different |  | 9.2.1 / A4.6
|  | ≤10% difference to calibration value | DF2 setting ____ _____ % different |  | 9.2.1 / A4.6
nvPM number analyzer operational checks: |  |  |  |  | 9.2.2
| Working fluid level |  |  |  | 9.2.2
| CPC saturator and condenser temperatures |  |  |  | 9.2.2
| Inlet flow rate |  |  |  | 9.2.2
Make-up flow rate |  |  | ________ slpm
Section 3 total flow rate |  | 25 ± 2 slpm | ________ slpm |  | 5.4
GL flow check
Transfer Part Cleanliness/Leakage check |  |  |  |  | 9.2.1 / A4.3
| Mass analyzer concentration | 30s average ≤1 µg/m3 | _______ µg/m3 |  | 9.2.1 / A4.3
| Number analyzer concentration | 30s average ≤2.0 p/cm3 | _______ p/cm3 |  | 9.2.1 / A4.3
Ambient nvPM measurement (if engine not in test cell) |  |  |  |  | 9.2.2 / A4.5
| Mass analyzer concentration |  | _______ µg/m3 |  | 9.2.2 / A4.5
| Number analyzer concentration (DF2 corrected) | 10 times >Cleanliness result | _______ p/cm3 |  | 9.2.2 / A4.5
| No data entry | Requirement | Recommendation or best practice

## OP During engine running
Test/Operation |  | Yes (tick) |  |  | ARP 6320 requirement reference
Prior to engine start:
Close Isolation valve 1
Begin backpurge through probe
Engine stable:
Turn backpurge off
New or cleaned section of nvPM sampling system? |  |  |  |  | 9.2.2
Open isolation valve to start nvPM sampling
| Expose nvPM sampling system for minimum 30 mins |  |  |  | 9.2.2
Ambient nvPM measurement (if testing inside a test cell) |  |  |  |  | 9.2.2 / A4.5
| Mass analyzer concentration |  | _____ µg/m3 |  | 9.2.2 / A4.5
| Number analyzer concentration (DF2 corrected) | 10 times >Cleanliness result | _______ p/cm3 |  | 9.2.2 / A4.5
nvPM measurement |  |  |  | Repeat for each nvPM measurement point
Engine on condition and stable (testpoint no.)
Open isolation valve to start nvPM sampling (if not already open)
| If P1 is at subatmospheric pressure, the P1 pressure control valve (and shut-off valve if installed) shall be closed |  |  |  | 9.2.2
| Calculate and report DF1_S using continuous GL (raw) and Makeup (diluted) CO2 concentration measurements |  |  |  | 9.2.2
| DF1_S is within 8 to 14 |  |  |  | 9.2.2
If DF1_S is >14 and system has capability, backpurge Diluter1 to attempt remove Diluter1 orifice blockage
| If P1 is above atmospheric pressure, the P1 pressure control valve shall be used to control DF1_S Depending on diluent flow control system, adjusting diluent flow could also be used to control DF1_S |  |  |  | 9.2.2
| Calculate and report DF1_S using continuous GL (raw) and Makeup (diluted) CO2 concentration measurements |  |  |  | 9.2.2
| DF1_S is controlled and maintained within 8 to 14 |  |  |  | 9.2.2
If DF1_s > 14, back purge through Diluter1 if option available
| Section 3 total flow rate 25 ± 2 slpm |  |  |  | 9.2.2
Is CPC in single count mode?
| Are nvPMmi, nvPMni and both GL and makeup CO2 signals stable? |  |  |  | 9.2.2
| Record 30 s average data |  |  |  | 9.2.2
Calculate 30 s average EImass and EI number
Calculate 2σ variation of nvPMmi, nvPMni, CO2_dil1, EIm and EIn
Monitor system and instrument parameters to check for failures
Perform nvPM cleanliness checks periodically during engine test
| Perform hourly zero and span checks for the gas analyzers (GL and Make-up flow) | Success criteria (<2% FS) |  |  | 9.2.2
Close the isolation valve and spill valve if necessary.
Prior to engine shutdown:
Ambient nvPM measurement (3 minutes average) |  |  |  |  | 9.2.2 / A4.5
| Mass analyzer concentration |  | _______ µg/m3 |  | 9.2.2 / A4.5
| Number analyzer concentration (DF2 corrected) | 10 times >Cleanliness result | _______ p/cm3 |  | 9.2.2 / A4.5
Backpurge through probe
| No data entry | Requirement | Recommendation or best practice

## OP After engine shutdown
Test/Operation |  | Criteria | Result | Pass/Done
Cleanliness check
| Mass analyzer concentration | 30s average ≤1 µg/m3 | _______ µg/m3
| Number analyzer concentration | 30s average ≤2.0 p/cm3 | _______ p/cm3
Stop system data logging
Back up all new data files
Power down all subsystems:
System heaters
pumps and flow controllers
Measurement analyzers (CO2, nvPMmi, VPR, nvPMni)
Data logging systems
Turn off compressed gas supplies, vent any excess pressure
DF1 Diluent
DF2 Diluent (VPR)
Compressed air (if required for nvPMmi)
CO2 span gases
| No data entry | Requirement | Recommendation or best practice

## Data System Parameters
ARP6320 | Parameter | Unit | ARP 6320 requirement reference
Table 9.1 | Section 2 Inlet pressure | Bar
| Diluent pressure | Bar
| GL temperature | K
| GL flow rate | slpm
| Raw CO2 concentration [CO2] | ppm
| Diluted CO2 concentration [CO2]dil1 | ppm
| 2σ diluted CO2 concentration after first stage | ppm
| DF1_S | -
| 2σ nvPM mass concentration | ug/m3
| 2σ nvPM number concentration | particles/cm3
| Mass analyzer (nvPMmi) sample temperature | K
| Mass analyzer (nvPMmi) sample pressure | Bar
| Number analyzer (nvPMni) sample temperature | K
| Number analyzer (nvPMni) sample pressure | Bar
| EImass | mg/kg-fuel | 3.3
| 2σ EIMass | mg/kg-fuel
| EInum | particles/kg-fuel | 3.3
| 2σ EInum | particles/kg-fuel
| average pre and post-test ambient air particle mass concentration readings | ug/m3 | 9.2.2
| 2σ ambient particle mass concentrations | ug/m3
| average pre and post-test ambient air particle number concentration readings | particles/cm3 | 9.2.2
| 2σ ambient particle number concentration | particles/cm3
| Cleanliness test - mass | ug/m3 | 9.2.2
| 2σ cleanliness mass | ug/m3
| Cleanliness test - number | particles/cm3 | 9.2.2
| 2σ cleanliness number | particles/cm3
| No data entry | Requirement | Recommendation or best practice

## REQ nvPM transport parameters
ARP6320 | System Parameter | Unit | Frequency of parameter change | Compliance check
A.2 | Probe inlet engine EGT / Section 1 temperature | K | per testpoint
| Section 2 Diluter1 inlet Temperature (T1) | K | per testpoint
| Section 2 Diluter1 temperature | K | per engine test
| Section 3 temperature | K | per engine test
| Section 4 cyclone separator temperature | K | per engine test
| Section 4 mass line temperature | K | per engine test
| Section 4 number line temperature | K | per engine test
| Section 1 & 2 flowrates (estimated) | slpm | per testpoint
| Section 3 and 4 total flow rate | slpm | per engine test
| Section 4 Mass analyzer flow rate | slpm | per engine test
| Section 4 VPR analyzer flow rate | slpm | per engine test
| Section 4 Makeup flow rate | slpm | per engine test
| Individual pipe IDs for section 1 | cm | per engine test
| Individual pipe IDs for sections 2 to 4 | cm | Fixed for specific system
| Individual pipe lengths for section 1 | cm | per engine test
| Individual pipe lengths for sections 2 to 4 | cm | Fixed for specific system
| total angle of sampling tube bends in Sections 1 and 3 | degrees | per engine test
| total angle of sampling tube bends in Sections 2, 4 and flexible tubing in Section 5 | degrees | Fixed for specific system
| cyclone d50 cutpoint | - | Fixed for specific system
| cyclone sharpness | % | Fixed for specific system
| diluter1 four penetration fractions Table 5.1 | - | Fixed for specific system
| VPR calibration four penetration values Table 7.1 per DF2 setting | - | per calibration
| CPC calibration two counting efficiencies | - | per calibration
| DF1 | - | per testpoint
| DF2 | - | per testpoint
| nvPM mass concentration | mg/m3 | per testpoint
| nvPM number concentration | particles/cm3 | per testpoint

## REQ Calibration&Maintenance
ARP6320 | Description | Component | Criteria | Requirements | Compliance check
9.1 | All instruments shall be maintained conforming to the manufacturer’s guidelines.
9.1.1 | Sampling system and Make-up Flow routine | Cyclone Separator | collection reservoir | shall be emptied and cleaned on an annual basis
|  | Flow rate controllers | NMI-traceable calibration | within 5% of Full Scale on an annual basis for: • Make-up flow • nvPMmi • VPR, • nvPMni
|  | Pressure transducers | NMI-traceable calibration | within 2% of Full Scale on an annual basis
|  | Diluter1 | Cleanliness | When DF1 approaches the upper range limit (14), clean Diluter1 as per manufacturer guidelines. As a minimum, Diluter1 sample inlet orifice nozzle shall be cleaned
9.1.2 / 6.3 / 6.1 | nvPMmi routine assessment | nvPMmi annual calibration by a competent laboratory | calibration document | slope of the linear regression between mass instrument and EC determined by TOT ASTM D6877-13 with temperature profile shown in Table 6.3
| nvPMmi compliance (performance certificate change) Performance demonstration after any hardware or software changes to the nvPMmi which affects data acquisition and processing by the manufacturer or competent laboratory |  | Range | 0 to 1000 mg/m3
|  |  | Resolution | 1 µg/m3 or better
|  |  | Interference | insensitivity to vPM
|  |  | Repeatability | 10 µg/m3
|  |  | Zero drift | 10 µg/m3/hr
|  |  | Linearity | 15 µg/m3
|  |  | LOD | 1 µg/m3
|  |  | Rise time | 2 sec
|  |  | Sample rate | 1 Hz
|  |  | Accuracy | 0.90 ≤ slope ≤ 1.10 Slope of the linear regression between mass instrument and EC determined by TOT ASTM D6877-13 with temperature profile shown in Table 6.3
|  |  | Applicability | ±16%
9.1.3 | VPR routine assessment | VPR annual calibration by a competent laboratory | Dilution Factor (DF2) | shall be measured or determined for each dilution setting of the VPR
|  |  | solid (non-volatile) particle penetrations at electrical mobility diameters | ≥30% at 15 nm ≥55% at 30 nm ≥65% at 50 nm ≥70% at 100 nm
|  |  | Volatile material | Tetracontane (CH3(CH2)38CH3) with purity >95% particles at 30 nm and inlet concentration
|  |  | Minimum inlet particle concentration | ≥10,000 particles/cm3
|  |  | Efficiency removal of volatile particles at an electrical mobility diameter | >99.5% removal of 30 nm particles
|  | VPR maintenance | Catalytic Stripper Replacement | As manufacturer guidelines
| VPR compliance (performance certificate change) VPR performance demonstration - after any hardware or software changes to the VPR which affects data acquisition and processing by the instrument manufacturer or competent laboratory | Sample Dilution Device | Dilution stages | one or more stages
|  |  | Heated section | 623 ± 15 K
|  | Diluted Sample | Level of dilution | ensure that sufficient dilution is applied to particle number concentration to maintain CPC in single particle count mode
|  |  | Temperature at CPC inlet | ensure temperature reduced to between 283 and 308 K (10 and 35°C)
|  |  | Pressure to CPC inlet | +/- 15 kPa of ambient pressure, and not exceed manufacturer maximum sample pressure specification
|  | CS | if included | diluent must include at least 10% O2 content
|  |  | if not used | place a heated dilution stage upstream which provides a sample temperature between 423 K and 638 K
dilutes by a factor > 8
|  | Penetration | solid (non-volatile) particle penetrations at electrical mobility diameters | ≥30% at 15 nm ≥55% at 30 nm ≥65% at 50 nm ≥70% at 100 nm
|  | Volatile Removal Efficiency | VRE | >99.5% removal of tetracontane (CH3(CH2)38CH3) with purity >95% particles at 30 nm and inlet concentration ≥10,000 particles/cm3
9.1.4 | nvPMni (CPC) routine assessment | CPC annual calibration by a competent laboratory | Counting accuracy | ±10% from 2000 particles/cm3 to upper threshold of single particle count mode against a traceable standard
|  |  | Counting efficiency curve at electrical moblity diameters | ≥50% at 10 nm and ≥90% at 15 nm determined with Emery Oil aerosol or another equivalent aerosol
|  | CPC maintenance | Wick replacement | replacement frequency as specified by manufacturer
|  |  | n-butanol working fluid replacement | replacement frequency as specified by manufacturer
| nvPMni compliance (performance certificate change) nvPMni (CPC) performance demonstration - after any hardware or software changes to the CPC which affects data acquisition and processing by the instrument manufacturer or competent laboratory | Performance Certificate | Working fluid | reagent grade n-butanol
|  |  | sample Flow | full flow operating conditions, no internal flow splitting
|  |  | Counting accuracy | ±10% from 2000 particles/cm3 to upper threshold of single particle count mode against a traceable standard
|  |  | resolution | ≥ 0.1 particles/cm3 at concentrations <100 particles/cm3
|  |  | Response | linear in single count mode
|  |  | Counting Mode | Single count mode only, photometric mode not allowed
|  |  | Data rate | ≥ 1.0 Hz
|  |  | t10-90 rise time | < 4s
|  |  | Coincidence | Maximum 10% coincidence correction. The coincidence correction function shall not use any algorithm to correct for or define the counting efficiency.
|  |  | Counting efficiency curve at electrical moblity diameters | ≥50% at 10 nm and ≥90% at 15 nm determined with Emery Oil aerosol or another equivalent aerosol
|  |  | Particle number concentration | reported at STP
9.1.5 | Gas Analyzer performance specification - CO2 | GL CO2 analyzer (ARP1256) | Range | Working ranges shall be selected based on the engine under test (the anticipated concentrations shall be within 20 to 95% FS)
|  |  | Zero gas purity | <100ppm CO2, <1ppm HC, <1ppm CO, <1ppm NOx
|  |  | Oxygen interference | < 0.1% of reading/percent of O2 concentration
|  | Make-up flow CO2 analyzer | Range | Approximately ten times lower than the CO2 analyzer used on the raw GL gaseous sample (the anticipated concentrations shall be within 20 to 95% FS)
|  |  | Zero gas purity | <10ppm CO2, <1ppm HC, <1ppm CO, <1ppm NOx
|  |  | Oxygen interference | Not required
|  | GL and Make-up CO2 analyzers (ARP1256) | Zero Drift | < 1% Full Scale in 1 hour
|  |  | Span Drift | < 1% Full Scale in 1 hour
|  |  | Linearity | < 1% Full Scale
|  |  | Noise | < 1% Full Scale
|  |  | Resolution | better than ±0.5% Full Scale
|  |  | Precision | < 1% Full Scale
|  |  | Response time to t90 | <10 seconds
| Gas Analyzer performance specification (ARP1256) - CO | CO | Zero Drift | < 1% Full Scale in 1 hour
|  |  | Span Drift | < 1% Full Scale in 1 hour
|  |  | Linearity | < 1% Full Scale
|  |  | Noise | < 1% Full Scale
|  |  | Resolution | better than ±0.5% Full Scale
|  |  | Precision | < 1% Full Scale
|  |  | Response time to t90 | <10 seconds
|  |  | Interference | Response of CO analyzer: ≤500 ppm per %Ethylene ≤2 ppm per %CO2 ≤2 ppm per %Water vapor
| Gas Analyzer performance specification (ARP1256) - HC | HC | Zero Drift | < 1% Full Scale in 1 hour
|  |  | Span Drift | < 1% Full Scale in 1 hour
|  |  | Linearity | < 1% Full Scale
|  |  | Noise | < 1% Full Scale
|  |  | Resolution | better than ±0.5% Full Scale
|  |  | Precision | < 1% Full Scale
|  |  | Response time to t90 | <10 seconds
|  |  | relative carbon response | Relative responses to n-hexane, propylene and toluene in air shall not differ by more than 5% from the relative response of propane in air
|  |  | O2 synergism | Response to Propane in 21% O2 shall not differ by more than 3% from the response to Propane in 10% O2
| Gas Analyzer performance specification (ARP1256) - NOx | NOx | Zero Drift | < 1% Full Scale in 1 hour
|  |  | Span Drift | < 1% Full Scale in 1 hour
|  |  | Linearity | < 1% Full Scale
|  |  | Noise | < 1% Full Scale
|  |  | Resolution | better than ±0.5% Full Scale
|  |  | Precision | < 1% Full Scale
|  |  | Response time to t90 | <10 seconds
|  |  | Interference | Response of NOx analyzer: ≤0.05 % of reading per %CO2 ≤0.1 % of reading per %Water vapor
|  |  | Converter efficiency | Greater than 90%

## REQ System General
General sampling system requirements
ARP6320 | Description | Component | Criteria | Requirements | Compliance check
5.1 | General | Probe inlet to measurement instrument inlet | Sampling line length | ≤ 35m
|  | Section 2 inlet to Section 4 outlet | Sampling line configuration | Straight-through as practical
|  |  | Bends | • only if necessary • radii ≥10 times the inside diameter of the line
|  |  | Fittings | • minimize number • stainless steel with a internal smooth bore
|  |  | Step-shoulders | • no forward facing >15% of the ID • changes >15% of ID only at splitter flow path interface
|  |  | Sample | Diluted within 8 m of probe tip
|  |  | Line material | where carbon loaded PTFE is used it should meet ISO8031 antistatic resistance limit
|  |  | Bulkhead union fittings | • keep number to a minimum • thermally insulated (no cold spots)
|  |  | Union interface | • heat throughout the union interface • if not practically possible, as a minimum, isolate the sample line from the interface surface and heat up to within 5cm of the interface surface and insulate thermally throughout
|  |  | Other sample line connection fittings | • Heat across the connection where possible • If not practically possible, heat the sample line up to within 5 cm of the next heated section and insulate thermally in-between

## REQ System Section1
Collection Part - Section 1 requirements
ARP6320 | Description | Component | Criteria | Requirements | Compliance check
5.2 | Section 1 | Probe/rake | Material | Stainless Steel or other non-reactive materials e.g Inconel 625 or Nimonic 75
|  | Probe/rake inlet to Splitter1 inlet | Probe placement and configuration | Probe shall provide a representative emission sample verified by means of detailed traverse measurement
|  |  | Probe axial location | perpendicular to nozzle shall not exceed 0.5 exit nozzle diameters from exit plane
|  |  | Number of Sampling Locations | ≥12 locations
|  |  | Total orifice area (multi-orifices probe) | at least 80% of the dynamic head pressure drop through the probe assembly is taken at the orifices
|  |  | Multiple sampling orifices | of equal diameter
|  |  | Sample Temperature | maintained at 433+/-15 K except for the section needed to cool the gas maintain ≥418 K if active probe cooling
|  |  | Length | ≤ 8 m
|  | Probe/rake outlet to Splitter1 inlet | Inner Diameter (ID) | 4 to 8.5 mm
|  |  | Material | stainless steel or carbon loaded PTFE

## REQ System Section2
Transfer Part - Section 2 requirements
ARP6320 | Description | Component | Criteria | Requirements | Compliance check
5.3 | Section2 - Splitter1 inlet to Diluter1 outlet | General | Length | ≤ 1m
|  |  | Material | Stainless steel or carbon-loaded electrically grounded PTFE
5.3.1 |  | Splitter 1 | Material | stainless steel
|  |  | General geometry | • Single or multiple sequential splitters • No forward facing step shoulders on the inner wall • nvPM flowpath straight-through and as short as practical • No change in ID from Splitter1 outlet to Diluter1 inlet • Flowpath to GL ID = 4 to 8.5 mm • Excess sample flow-path cross sectional area ≥ total inlet area of the probe tips • Excess sample line ID ≥ Splitter1 outlet ID
|  |  | Split angles | • as acute as practical
• ≤ 35°
|  |  | Temperature | maintained ≥418 K
5.3.2 |  | Isolation Valve 1 | Geometry | • full bore • No forward facing step shoulders > 15% to sample line ID)
|  |  | seal | • seals: dry and heat resistant to 448K (175°C)
5.3.3 |  | Diluter1 inlet interface | Temperature | wall temperature heated to within 5 cm of the Diluter1 mixing plane, T1 ≥418 K
5.3.4 |  | Diluter1 | Geometry | Inlet ID ≥ 7.59mm
|  |  | Type | ejector-type to provide positive pressure and consistent sample flow to Section 4
|  |  | Vent | open to ambient
|  |  | Temperature | Diluter1 body = 333 ± 15K (60 ± 15°C)
|  |  | Dilution Factor | 8 to 14
|  |  | Inlet sample pressure sensitivity | unmixed sampling - DF1 controlled to within the range 8 to 14 (for a Diluter1 inlet pressure range of -5,500 to +5,500 Pa (-55 to +55 mbar) relative to ambient)
Mixed flow sampling - DF1 controlled <14
|  |  | Penetration efficiency | same methodology as utilised for VPR (A6.1.4) with the required penetrations at four sizes (Table 5.1)
|  |  | Diluent | • Nitrogen or Air • HEPA filtered • contain <10ppm CO2 • diluent flow shall be controlled • heated (to provide a diluted PM sample temperature of 333±15 K (60±15 °C) at the outlet of Section 2)
5.3.6 |  | Excess sample line | Pressure | P1 maintained near local ambient air pressure
|  |  | Pressure control valve | • sufficient internal area to maintain P1 • capable of operating at 10,000Pa (-100mbar) relative to ambient • when valve is closed, it shall be capable of holding a vacuum pressure of -75kPa relative to ambient
5.3.5 / ARP1256 |  | Gas Line | Material | Stainless steel or carbon-loaded electrically grounded PTFE
|  |  | Length & Flowrate | Such that the sample transport time from sampling probe to analyzer inlet does not exceed 10 seconds
|  |  | ID | 4.0 to 8.5 mm
|  |  | Temperature for CO2 measurement only | 338 ± 15 K (65 ± 15 °C) upto dryer then 10 K above sample dewpoint
|  |  | Temperature for CO2, CO, HC and NOx measurement | 433 ± 15 K (160 ± 15 °C) then: 433 ± 15 K (160 ± 15 °C) for HC 338 ± 15 K (65 ± 15 °C) for NOx 338 ± 15 K (65 ± 15 °C) for CO/CO2 upto dryer, then 10 K above sample dewpoint
| Gas Line measurements from Section 2 | Gas Line Analyzers | CO2 analyzer (and CO, HC & NOx analyzers if used) | ARP1256 performance specifications (see Cal&Maintenance Requirements sheet)
|  |  | Dilution Factor | DF1 shall be calculated using undiluted and diluted CO2 measurements simultaneously with nvPM measurement

## REQ System Section 3
Transfer Part - Section 3 requirements
ARP6320 | System section | Component | Criteria | Requirements | Compliance check
5.4 | Section 3 | Diluter 1 outlet to Cyclone Separator inlet | Material | carbon-loaded, electrically grounded PTFE
|  |  | ID | 7.59 to 8.15 mm
|  |  | Length | 24.5±0.5 m
|  |  | Segments | maximum 3 segments
|  |  | Sampling line temperature | 333 ± 15 K (60 ± 15 °C)
|  |  | Coiled bend | radii > 0.5 m
|  |  | Flow rate | 25±2 slpm monitored by three measurements: nvPMmi, VPR & make-up flow rates

## REQ System Section 4
Transfer Part - Section 4 requirements
ARP6320 | System section | Component | Criteria | Requirements | Compliance check
5.5 | Section 4 - Cyclone Separator inlet to instrument inlets | General | Length | short as practical and ≤ 3 m
|  |  | Material | Stainless steel or carbon loaded electrically grounded PTFE
|  |  | Temperature | 333 ± 15 K (60 ± 15°C)
|  | Cyclone Separator | Material | Stainless steel
|  |  | Temperature | 333 ± 15 K (60 ± 15°C)
|  |  | Cut-point (at 25slpm) | D50 = 1.0 ± 0.1 µm
|  |  | Sharpness (at 25 slpm) | (D16/D84)0.5 ≤ 1.25
|  |  | Pressure-drop | Δp ≤ 2000 Pa (20 mbar)
|  |  | inlet & outlet ID | difference with sample line ID <15%
|  | Splitter2 | Material | Stainless steel
|  |  | General geometry | • single or multiple sequential splitters allowed • no forward facing step shoulders on the inner wall • nvPM flowpath as short as practical
|  |  | Split angles | • as acute as possible • ≤ 35°
|  |  | Flow paths split | • nvPMmi • volatile removal device (for nvPMni) • make-up flow
|  |  | Specific geometry | ≤15% difference in ID from Splitter2 outlet to nvPMmi inlet
≤15% difference in ID from Splitter2 outlet to VPR inlet
|  |  | Temperature | T3 = 333 ± 15 K (60 ± 15 °C)
• thermocouple placed in make-up flow-path at the outlet of Splitter2
|  | Instrument inlet lines | ID | ≤15% difference in ID between the sampling line and instrument inlet

## REQ System Section 5
nvPM measurement - Section 5 requirements
ARP6320 | System section | Component | Criteria | Requirements | Compliance check
5.6.1 / ARP1256 | Section 5 - Make-up flow line | Pump and flow controller | Flowrate | Capable of maintaining a constant total sample flowrate of 25 ± 2 slpm upto 10,000 Pa (100 mbar) below ambient, through Section 3
|  | CO2 analyzer | Location | after flow controller
|  |  | range | Approximately ten times lower than the CO2 analyzer used on the raw GL gaseous sample (the anticipated concentrations shall be within 20 to 95% FS)
|  |  | zero gas purity | <10ppm CO2, <1ppm HC, <1ppm CO, <1ppm NOx
|  |  | Oxygen interference | Not required
|  |  | Zero Drift | < 1% Full Scale in 1 hour
|  |  | Span Drift | < 1% Full Scale in 1 hour
|  |  | Linearity | < 1% Full Scale
|  |  | Noise | < 1% Full Scale
|  |  | Resolution | better than ±0.5% Full Scale
|  |  | Precision | < 1% Full Scale
|  |  | Response time to t90 | <10 seconds

## REQ nvPMmi
nvPM mass instrument requirements
ARP6320 | Requirement | Component | Criteria | Requirements | Compliance check
6.1 | nvPMmi performance specification demonstration - manufacturer or competent laboratory | nvPMmi | Range | 0 to 1000 mg/m3
|  |  | Resolution | 1 µg/m3 or better
|  |  | Interference | insensitivity to vPM
6.2 |  |  | Repeatability | 10 µg/m3
|  |  | Zero drift | 10 µg/m3/hr
|  |  | Linearity | 15 µg/m3
|  |  | Limit of Detection | 1 µg/m3
|  |  | Rise time | 2 sec
|  |  | Sample rate | 1 Hz
|  |  | Accuracy | 0.90 ≤ slope ≤ 1.10 Slope of the linear regression between mass instrument and EC determined by TOT ASTM D6877-13 with temperature profile shown in Table 6.3
|  |  | Applicability | ±16%
6.3 / 9.1.2 | nvPMmi calibration - manufacturer or competent laboratory | nvPMmi | calibration document | Slope of the linear regression between mass instrument and EC determined by TOT ASTM D6877-13 with temperature profile shown in Table 6.3

## REQ VPR
VPR instrument requirements
ARP6320 | Requirement | Component | Criteria | Requirements | Compliance check
7.1 | VPR performance specification demonstration - manufacturer or competent laboratory | Sample Dilution Device | Dilution stages | one or more stages
|  |  | Heated section | 623 ± 15 K
|  | Diluted Sample | Level of dilution | ensure that sufficient dilution is applied to particle number concentration to maintain CPC in single particle count mode
|  |  | Diluent cleanliness | HEPA filtered diluent
|  |  | Temperature at CPC inlet | ensure temperature reduced to between 283 and 308 K (10 and 35°C)
|  |  | Pressure to CPC inlet | +/- 15 kPa of ambient pressure, and not exceed manufacturer maximum sample pressure specification
|  | Catalytic Stripper | if included | diluent must include at least 10% O2 content
|  |  | if not included | place a heated dilution stage upstream which
provides a sample temperature between 423 K and 638 K
dilutes by a factor > 8
|  | Penetration | solid (non-volatile) particle penetrations at electrical mobility diameter | • ≥30% at 15 nm
• ≥55% at 30 nm
• ≥65% at 50 nm
• ≥70% at 100 nm
|  | Volatile Removal Efficiency | Volatile material | Tetracontane (CH3(CH2)38CH3) with purity >95% particles at 30 nm and inlet concentration
|  |  | Minimum inlet particle concentration | ≥10,000 particles/cm3
|  |  | Efficiency removal of volatile particles at an electrical mobility diameter | >99.5% removal of 30 nm particles
A6.1 | VPR calibration - manufacturer or competent laboratory | VPR | Dilution Factor (DF2) | shall be measured or determined for each dilution setting of the VPR
|  |  | solid (non-volatile) particle penetrations at electrical mobility diameters | ≥30% at 15 nm ≥55% at 30 nm ≥65% at 50 nm ≥70% at 100 nm
|  |  | Volatile material | Tetracontane (CH3(CH2)38CH3) with purity >95% particles at 30 nm and inlet concentration
|  |  | Minimum inlet particle concentration | ≥10,000 particles/cm3
|  |  | Efficiency removal of volatile particles at an electrical mobility diameter | >99.5% removal of 30 nm particles
7.3 | VPR to CPC interface | Sample line to CPC | Material | electrically conductive material
|  |  | ID | ≥4 mm
|  |  | Residence time | ≤ 0.8 s
|  |  | CPC inlet pressure (if CPC not reporting measured concentrations at STP) | Pressure measurement accuracy better than 2%
A4.6 | DF2 check determination equipment | Diluent | Cleanliness and composition | HEPA filtered air or N2 O2 ≥10% (if Catalytic Stripper used)
|  | CO2 analyzer for option (1) | ARP1256 compliant | See Cal & Maintenance Requirements sheet
|  |  | Range | FS 30 to 300 ppm measured concentrations as low as 10ppm
|  |  | Diluent impurity | CO2 <0.1 ppm dependent on DF2 setting
|  | CO2 analyzer for option (2) | ARP1256 compliant | See Cal & Maintenance Requirements sheet
|  |  | Range | FS 0.5 to 5 % dependent on DF2 setting
|  |  | Diluent impurity | CO2 <10 ppm

## REQ nvPMni
nvPM number instrument requirements
ARP6320 | Requirement | Component | Criteria | Requirements | Compliance check
7.2 | nvPMni performance specification demonstration - manufacturer or competent laboratory | nvPMni | Working fluid | reagent grade n-butanol
|  |  | Instrument sample flow | full sample flow operating conditions, no internal sample flow splitting
|  |  | Counting accuracy | ±10% from 2000 particles/cm3 to upper threshold of single particle count mode against a traceable standard
|  |  | Resolution | ≥ 0.1 particles/cm3 at concentrations <100 particles/cm3
|  |  | Response | linear in single count mode
|  |  | Counting Mode | Single count mode only, photometric mode not allowed
|  |  | Data rate | ≥ 1.0 Hz
|  |  | t10-90 rise time | < 4s
|  |  | Coincidence | maximum 10% coincidence correction. The coincidence correction function shall not use any algorithm to correct for or define the counting efficiency.
|  |  | Counting efficiency curve at electrical monility diameters | ≥50% at 10 nm and ≥90% at 15 nm determined with Emery Oil aerosol or another equivalent aerosol
|  |  | Particle number concentration | reported at STP
A6.2 | nvPMni calibration - manufacturer or competent laboratory routine check | CPC | Counting accuracy | ±10% from 2000 particles/cm3 to upper threshold of single particle count mode against a traceable standard
|  |  | Counting efficiency curve at electrical moblity diameters | ≥50% at 10 nm and ≥90% at 15 nm determined with Emery Oil aerosol or another equivalent aerosol